## Slide 1: WNR Weekly Meeting
November 5, 2015

## Slide 2: Agenda
Power Feasibility of Intan Chip and BLE (If required)
Transmission Constraints: Sampling Rate and Precision
Finalizing Intan Chip purchase and capabilities
BLE CC2650 Development Updates
Nordic nRF52 Development Updates
TCP/UDP transmission protocols
Compression/Huffman Encoding Preliminary Results
EEG Data and Obtaining Electrodes

## Slide 3: Power Feasibility

## Slide 4: Power Feasibility: AFE - Intan RHD2000
Amplifier:
  Baseline = 200 uA
  7.6 uA/kHz upper cutoff frequency/channel
ADC:
  Baseline = 510 uA
  2.14 uA/kHz/channel
Supply Voltage and Auxiliary Inputs:
  40 uA

## Slide 5: Power Feasibility: AFE - Intan RHD2000
Amplifier Current:
  200 uA + 16 * 1* 7.6 uA/Khz/channel = 321.6 uA
ADC Current:
  510 uA + 16 * 1 * 2.14 uA/kHz/channel = 544.24 uA
Supply Voltage and Auxiliary Inputs:
  40 uA
Total Power Consumption with 3.3V Supply
  3.3 * (321.6 + 544.24 + 40) * (10^(-6)) = 2.99 mW
  To run for 24 hours: (321.6 + 544.24 + 40) uA * 24 hours = 21.74 mAh

## Slide 6: Power Feasibility: BLE Transmission - TI CC2650
Transmit at +0 dBm:
  6.1 mA
  Power Consumption with 3.3 V power supply: 3.3V * 6.1 mA * (10 ^ -3) = 20.13 mW
  To run for 24 hours: 6.1 mA * 24 hours = 146.4 mAh
Transmit at +5 dBm:
  9.1 mA
  Power Consumption with 3.3V power supply: 3.3V * 9.1 mA * (10^-3) = 30.03 mW
  To run for 24 hours: 9.1 mA * 24 hours = 218.4 mAh

## Slide 7: Power Feasibility: Battery
Zinc Air p13:
  diameter: 7.9 mm
  height: 5.4 mm
  Typical capacity 300 mAh

## Slide 8: Power Feasibility: Total System
AFE: 21.74 mAh
Transmitting at +0 dBm: 146.4 mAh
  Total Power Consumption = 21.74 + 146.4 = 168.14 mAh
Transmitting at +5 dBm: 218.4 mAh
  Total Power Consumption = 21.74 + 218.4 = 240.14 mAh
Using the Zinc-Air p13 battery (300 mAh) is enough to power both types of transmission with enough power for compression and other processing

## Slide 9: Transmission Bandwidth

## Slide 10: Transmission Bandwidth
BLE:
  Theoretical Limit: 1 Mb/s
  Usable Bandwidth: ~750 kb/s
  Aim for: ~500 kb/s
Given Specifications:
  Channel Sampling Rate: 1 kS/s
  Number of Channels per electrode: 16
  Number of electrodes per array: 16
  Sampling Precision: 8-bits
    Commercial products, such as Neuropace use 10-bit precision
  Bandwidth Required:
    1 (kS/s)/(channel) * 16 channels/electrode * 16 electrodes * 8-bits/Sample = 2 Mb/s

## Slide 11: Transmission Bandwidth
Reduce transmission rate and sampling rate to ~250 Hz
  Bandwidth Required with no Compression:
    1/4 (kS/s)/(channel) * 16 channels/electrode * 16 electrodes * 8-bits/Sample = 512 kb/s
Compression (even with just Huffman Encoding) has the possibility of compressing the data by ~40%

## Slide 12: Tasks Remaining
Measure power consumption of Nordic nRF52 board while transmitting
Interface with and measure the power consumption of Intan Chip after it arrives.
Combine individual components into the system.